## Ref Values
English XL w/Seq Tile E |  | English XL w/Seq Tile F |  | English XL w/Seq Tile G
xdbar | 0.14661 | xdbar | 0.1951 | xdbar | 0.29741
sr | 0.010615 | sr | 0.0075465 | sr | 0.012584
r | 0.029721 | r | 0.02113 | r | 0.035235
sR | 0.024154 | sR | 0.016252 | sR | 0.030903
R | 0.067632 | R | 0.045506 | R | 0.08653
R* | 0.050611 | R* | 0.034053 | R* | 0.064753
low limit | 0.095996 | low limit | 0.16105 | low limit | 0.23266
high limit | 0.19722 | high limit | 0.22915 | high limit | 0.36216
# TR < Low Lim | 4 | # TR < Low Lim | 3 | # TR < Low Lim | 2
# TR > High Lim | 4 | # TR > High Lim | 6 | # TR > High Lim | 6
F2508-23 9.1.1 Rank Order
Tile E test result < Tile F test result < Tile G test result
F2508-23 9.1.2 Differentiation
(Tile F test result) - (Tile E test result) > Larger of (Tile E r) or (Tile F r)
(Tile G test result) - (Tile F test result) > Larger of (Tile F r) or (Tile G r)
F2508-23 9.1.3 Within Established Limits
(Tile E low limit) < Tile E test result < (Tile E high limit)
(Tile F low limit) < Tile F test result < (Tile F high limit)
(Tile G low limit) < Tile G test result < (Tile G high limit)

## F2508-23 Val XLS
ASTM F2508-23 Validation Check |  |  |  |  | Validation performed in accordance with F2508-23
Operator |  |  | P. Widas |  | * Only applicable to this User, Instrument, Test Foot
Test Address |  |  | 321-C Johnstown Rd, Chesapeake, VA 23322
Company |  |  | Excel Tribometers, LLC
Contact Info |  |  | pwidas@exceltribometers.com, 757-897-2853
Source of Ref Surfaces |  |  | ASTM 2024 Sticker
Date Acquired |  |  | 2026-01-27
Validation Test Date |  |  | 2026-02-16 |  | Previous Validation Date |  | NA
Validation Temperature |  |  | 70.7°F
Validation Humidity |  |  | 46%
Walkway Tribometer Supplier |  |  | Excel Tribometers, LLC
Walkway Tribometer Model |  |  | English XL VIT w/Sequencer
Walkway Tribometer Serial Number |  |  | 2602XXX |  | Sander SN |  | 13X
Test Foot Number |  |  | 108XX
Test Foot Material |  |  | Neolite
Test Foot Age (Smithers Manufacture) |  |  | 2025-08-07 |  | Test Foot Age (Mounted) |  | 2026-01-21
Test Foot Preparation Procedure |  |  | Calibrate standard XL VIT test foot per the XL User Guide
Test Foot Dimensions (Diameter, in) |  |  | 1.25
Operating Instructions Version |  |  | v4 3/21/16
Calibration Tile Serial Number: |  |  | 6XX |  | Tile Certified Value: | 0.23±0.03
Test Foot Verification - Distilled Water
N | N | N | E | S | W |  | AVG (last 4)
0.27 | 0.25 | 0.24 | 0.22 | 0.22 | 0.23 |  | 0.2275
Reference Material E - 0.05% SLS Solution
N | N | N | E | S | W |  | AVG (last 4)
0.21 | 0.18 | 0.17 | 0.17 | 0.17 | 0.16 |  | 0.1675
Reference Material F - 0.05% SLS Solution
N | N | N | E | S | W |  | AVG (last 4)
0.27 | 0.22 | 0.22 | 0.21 | 0.2 | 0.2 |  | 0.2075
Reference Material G - 0.05% SLS Solution
N | N | N | E | S | W |  | AVG (last 4)
0.32 | 0.3 | 0.29 | 0.31 | 0.31 | 0.31 |  | 0.305
Tile E Result | 0.1675 |  | Tile F Result | 0.2075 |  | Tile G Result | 0.305
9.1.1 Rank Order: E < F < G |  |  | E<F | Pass |  | F<G | Pass
9.1.2.1 Differentiation
Diff F-E | 0.04 |  | Diff G-F | 0.0975
Diff > Tile E r | Pass |  | Diff > Tile F r | Pass
Diff > Tile F r | Pass |  | Diff > Tile G r | Pass
9.1.3 Within Established Limits
E Low Limit | Pass |  | F Low Limit | Pass |  | G Low Limit | Pass
E High Limit | Pass |  | F High Limit | Pass |  | G High Limit | Pass
F3539 ILS References
English XL w/Seq Tile E |  |  | English XL w/Seq Tile F |  |  | English XL w/Seq Tile G
xdbar | 0.14661 |  | xdbar | 0.1951 |  | xdbar | 0.29741
sr | 0.010615 |  | sr | 0.0075465 |  | sr | 0.012584
r | 0.029721 |  | r | 0.02113 |  | r | 0.035235
sR | 0.024154 |  | sR | 0.016252 |  | sR | 0.030903
R | 0.067632 |  | R | 0.045506 |  | R | 0.08653
R* | 0.050611 |  | R* | 0.034053 |  | R* | 0.064753
low limit | 0.095996 |  | low limit | 0.16105 |  | low limit | 0.23266
high limit | 0.19722 |  | high limit | 0.22915 |  | high limit | 0.36216

## Cleaning
Per F2508-23 Section 8.3.1
8.3.1 Cleaning:
8.3.1.1 No surface treatment except as specified in this section is permitted.
8.3.1.2 Disperse at least 25 mL of 0.05 % SLS solution onto the RM surface. Gently wipe with a clean microfiber cloth, in a circular pattern, the entire RM test surface for a minimum of 10-s. Reapply 25 mL of 0.05 % SLS and repeat the minimum 10-s of wiping with the same microfiber cloth. Ensure that the
0.05 % SLS solution has been applied to the entire RM test surface.
NOTE 1—To clean the microfiber cloth after use, run a basin of cool or warm water, insert the cloth and agitate it by hand. Soak the cloth for a minimum of 15 min then agitate it again by hand. Remove it from the water basin and thoroughly rinse under cool or warm running water. Wring out as much remaining liquid as possible by hand, then hang dry.
8.3.1.3 Rinse the surface thoroughly with distilled or deionized water, ensuring that no visible suds or residues remain.
8.3.1.4 Prepare an ethanol solution containing equal parts denatured ethanol and distilled water.
8.3.1.5 Disperse at least 25 mL of ethanol solution onto the RM surface. With a clean microfiber cloth different from that used in 8.3.1.2, wipe the RM test surface for 10-s.
8.3.1.6 Dry the RM with dry and oil-free compressed air or air dry if compressed air is not available. Any visible contamination remaining after this step will disqualify the RM for use.
8.3.1.7 Ensure that handling of the RM does not introduce contaminants to the test surfaces, including exposing the test surfaces to contact by human skin.

## Test Procedure
Summary of Procedure:
1. Labs will prepare the test foot
2. Labs will perform test foot calibration (verification) using distilled water
3. Labs will fully dry the test foot and recondition by sanding prior to testing each RM
4. Each RM will be cleaned according to F2508-23 prior to testing.
5. Labs will conduct testing on the RMs using 0.05% SLS Solution as the contaminant
6. Labs will conduct 4 cleaning slips, and 2 preliminary readings that are considered cleaning
7. Labs will conduct 4 subsequent orthogonal readings in the center of the RM (test observations)
8. Test observations (4 last readings) will be averaged to obtain a test result
9. Labs will fully dry and prepare the test foot by sanding in preparation for the next RM
10. Labs will repeat this process to produce a test result on each RM.